BASIC CT PROTOCOLS

Nasal cavity

Dolichocephalic nose

- Pre and post contrast

- Axial

- Sharp or detail filter

- High edge enhancement

- 35 mm thickness

- 5 mm interval

- scan from nasal planum through retropharyngeal lymph nodes

- If lesion affects cribiform/ethmoids, reposition patient in dorsal recumbency with nose pointing to ceiling.  Scan this dorsal plane of at 1mm thickness/ 1mm interval with bone filter.

Positioning:

- Sternal recumbency

- Head first

- Head should be straight!

- Plan scans on DV scout view

Small canine / feline nose

- Pre and post contrast

- Axial

- Sharp or detail filter

- High edge enhancement

- 3mm thickness

- 3mm interval

- If lesion affects cribiform/ethmoids, reposition patient in dorsal recumbency with nose pointing to ceiling.  Scan this dorsal plane of at 1mm thickness/ 1mm interval with bone filter.

Positioning:

- Sternal recumbency

- Head first

- Plan scans on DV scout view

Orbital region

Orbit, axial plane scan

- Pre and post contrast

- Axial scan

- Detail or sharp filter

- No edge enhancement

- 2 mm slice thickness

- 2 mm slice interval

Positioning:

- Sternal recumbency

- Head first

- Plan scans on DV or lateral scout view

Brain

Large canine / small canine & feline brain scan

- Pre and post contrast

- Axial

- Detail filter

- No edge enhancement

- 5 mm slice thickness / 3 mm for small patients

- 5 mm slice interval /3 mm for small patients

Positioning:

- Sternal recumbency, with the forelimbs pulled caudally alongside the chest

- Head first

Plan scans on DV scout view

Spine

Spine – intervertebral disc space

- Pre contrast (if lesion seen on survey radiographs) or post myelogram

- Axial scan

- Detail filter

- No edge enhancement

- 1 mm slice thickness

- 1 mm slice interval

Positioning:

- Dorsal recumbency

- Head or tail first

- Gantry angle to align with intervertebral disc space(s).  Change angle for each disk space if needed.

- Plan scans on lateral scout view

Spine, high-resolution bone scan (i.e. bone tumors, diskospondylitis, fracture)

- Pre Contrast

- Axial scan

- Bone filter

- High edge enhancement

- 2 mm slice thickness

- 1 mm slice interval

- Positioning:

- Dorsal recumbency

- Head or tail first

- Gantry rotation to align with intervertebral disc space(s)

- Plan scans on lateral scout view

General spine, pre contrast helical scan

- Helical

- Detail filter

- No edge enhancement

- 5 mm slice thickness

- 5 mm slice interval

- 1.3 or 1.5 pitch

- Once lesion is identified, repeat with smaller slice thickness and interval with axial images both pre and post contrast

Positioning:

- Dorsal recumbency

- Head or tail first

- Gantry rotation to align with intervertebral disc space(s)

- Plan scans on lateral scout view

Thorax/mediastinum

Mediastinal / Cardiovascular helical scan

- Pre and Post Contrast

- Helical

- Soft tissue filter

- No edge enhancement

- 5 mm slice thickness

- 1.5 pitch

Positioning:

- Sternal or dorsal recumbency

- Head first

- Plan scans on DV or lateral scout view

Lung

- No or Pre Contrast

- Helical

- Sharp/detail filter

- High edge enhancement

- 3-5 mm slice thickness depending on patient size

- 1.5 pitch

- Hyperventilate prior to scan

- Manual breath hold during scan

Positioning:

- Sternal recumbency

- Head first

Plan scans on DV or lateral scout view

##### Abdomen (especially renal and adrenal) and pelvis

Abdomen, large dog, helical scan

- Pre and post contrast

- Helical

- Detail filter

- No edge enhancement

- 5 mm slice thickness

- 1.5 pitch

Positioning:

- Sternal or dorsal recumbency

- Head or tail first

- Plan scans on lateral scout view

Abdomen, small canine / feline, helical scan

- Pre and post contrast

- Helical

- Detail filter

- No edge enhancement

- 3 mm slice thickness

- 1.5 pitch

Positioning:

- Sternal or dorsal recumbency

- Head first

- Plan scans on lateral scout view

Dual phase CT Portography

- Non Contrast Portion of Scan

- Dorsal recumbency

- 18-20 G cephalic catheter

- Pressure injector at 5ml/sec (if you have one, otherwise begin scan at end of rapid injection)

- Hyperventilate prior to scan

- Pilot image like VD radiograph

- Non-contrast survey

- KvP 130, MA 200

- Label as Non-contrast

- Cranial aspect of diaphragm to caudal L5

- 3mm collimation with 1.7 pitch or 5mm with 1.4 pitch

- Identify slice where porta hepatis is seen

- Dual Phase/Helical Portion of Scan

- Hyperventilate prior to scan +/- breath hold during scan

- Label as + Contrast

- Inject 814mg I/kg

- Arterial phase

  - Start scan at time of injection.  There is an inherent 6-8 second delay of scanner to start

  - Plan with “load and go”—see below

  - Scan from porta hepatis to cranial aspect of diaphragm

- Portal phase

  - Begin as soon as arterial phase acquisition is complete

  - Scan from cranial aspect of diaphragm to caudal L5

  - Plan with “load and go” (plan study ->Load and go (arterial phase)->plan additional study (portal phase)

**Arterial and portal phases should not be separated by any break.  Scanner should be programmed for both and scans should take place back to back.

##### Musculoskeletal system and superficial soft tissues

Elbows, high-resolution bone scan: scanning protocol

- No Contrast

- Helical

- Bone filter

- High edge enhancement

- 1 mm slice thickness

- 1.0-1.3 pitch

- scan from olecranon process to distal to coronoid process of ulna

- Reconstruct the images to center on each elbow separately with smaller image size based on size of patient (usually 10cm)

Positioning for elbows (both elbows on one scan):

- Head-first

- Dorsal recumbency in a trough

- Elbows positioned at same level

- Extend both legs & fix them (sandbags, tape)

- Plan scans on DV scout view